## Slide 1: Development of the optical modules for IceCube-Gen2 Nobuhiro Shimizu (Chiba University) Shingakujutsu workshop A03 2022 March 7th

## Slide 2: The IceCube and future upgrades
IceCube Gen-1
a cubic kilometer neutrino telescope deployed in Antarctica
5160 Digital Optical module (DOM) to detect Cherenkov photons produced by secondary particles from neutrino interaction
IceCube Upgrade
Deploy 700 new optical modules from 2023
IceCube Gen-2
Deploy (10,000) new optical modules from 2027
Digital Optical Module (DOM)
for Gen-1

## Slide 3: New optical modules for Gen2
Requirements of new optical modules
  Volume of Gen2  x 8
  Sparse layout: 125 m  240 m
  low cost per optical sensitivity
2023
2027
Reduce a diameter to save cost of drilling
33 cm
30.5 cm
Gen1 DOM
D-Egg
30 cm
LOM
Long optical module (LOM)
House as many PMTs as possible
① Use a low noise elongated glass vessel.
② House 18 four inch PMTs. ③ Adopt “gel pads” to efficiently lead photons to PMTs

## Slide 4: Concept of the design of LOM gel pad scheme
A gel pad
Inspired by HK’s multi-PMT module. The UV-transparent silicone elastomer is casted to a conical shape.
Works as
1. an optical interface and
2. a photon collector thanks to internal reflection.
A golden color of the photocathode
 implies the enhancement of effective area

## Slide 5: Improvement of sensitivity
②
④
①
④ Gel pads
③ Fill gel
② No optical interface
photons
The D-Egg already has x2.8 higher sensitivity than IceCube-Gen1 DOM.
Optical interface is necessary
To fill UV-transparent gel may be good but increases the cost and weight.
With the gel pad approach, we further expect x1.5 improvement of sensitivity compared to D-Egg with cost-effective way.
① D-Egg
③

## Slide 6: Development of the components

## Slide 7: Pressure-resistant glass of LOM
Requirements on the glass
Pressure-resistant up to 70 MPa while small material to reduce price
Transparency to UV photons
Low noise due to RI sources
1.4 mm
Horizontal variation in 70 MPa
0.7 mm
Chart: 
| K | U238 | U235 | Th
 | 2.55 | 0.175 | 0 | 0.14
Bq/kg
RI decays in the glass
Transparency of glass
Thickness:
20 mm (D-Egg)  16.5 (LOM)

## Slide 8: Pressure test
JAMSTEC (2020 March 3rd)
Attach strain sensors inside the glass and measured the strain of glass
Test: 50 MPa and 70 MPa runs
  No notable change observed in the glass
Sample 1 (50 MPa run)
Sample 2 (70 MPa run)
Location of sensor | Measured strain at 70 MPa (mST) | Simulated strain at 70 MPa (mST)
Bottom | 5.0 | 5.8
Curved | 6.2 | 5.8
Equator | 9.2 | 9.3
Equator (max) | 9.7 | 9.6
Hole | 4.8 | 4.3
curved
hole
bottom
equator
equator (max)
 Reasonable agreement between the measured and simulated strains.

## Slide 9: New 4 inch PMT
ns
4 inch PMTs were not in the existing lineup of Hamamatsu
Requirements for the new PMTs
  as high QE as possible (typically better than 25%)
  TTS 6 ns
  height  shorter than 106 mm
  diameter  larger than 102 mm
Start the development from autumn in 2020
1st version arrived at summer in 2021
< 106 mm
> 102 mm
The new Hamamatsu PMT satisfied our basic requirements.

## Slide 10: Readout of signals
digitization
(wuBase)
x18
MUX
board
“mini”
main
board
surface
DAQ
wuBase
HV generation with Cockcroft-Walton circuit
Digitize the waveform of signal with 60 MSPS (high and low gain channels)
Super low power consumption (145 mW/PMT)
Digitized waveform
MUX board
”fanout” of 9 PMT signals
programable coincidence of multi-PMT hits
mini main board
communicate with surface DAQ
power management ( 48 V  low voltages)
storage of hit data until trigger signal comes
x2 (upper
and lower)

## Slide 11: Internal design of LOM
Structure
SUS plates to hold PMTs
center shaft to fix SUS plate
springs to push PMTs outwards

## Slide 12: Casting of gel pads
A vacuum forming technique to make a plastic disposable mold.
After mixing two component silicone, pour liquid gel into the mold.
After waiting a few days to cure
We have almost established a method to cast a beautiful gel pad on a PMT.
 Further study needed to make the procedure efficient/systematic.

## Slide 13: Installation procedure of the gel pads
①
②
③
④
⑤
⑥
Preparation of a premade gel pad casted on a PMT
Dispense liquid gel
Install a layer (four PMTs)
Remove bubbles in vacuum
Release vacuum and wait for curing

## Slide 14: Summary
IceCube Gen-2 plans to deploy O(10,000) new optical sensors. To accommodate the larger volume, the sensors are required to have higher sensitivity.
A new sensor LOM houses 18 four-inch PMTs in the glass, and expected to improve the sensitivity by a factor of 1.5 compared to D-Egg.
Development of each component is under way
  Pressure resistant glass tolerated up to 70 MPa
  New 4 inch PMTs were developed and satisfied our basic requirements
  Electronics are finalized
Establish of the assembly procedure is now ongoing.
  Efficient and systematic casting of gel pads on PMTs
  A prototype assembly using 3 inch PMTs + D-Egg glass was performed.

## Slide 15: Backup

## Slide 17: Realistic 3D model

## Slide 18: Installation procedure
① Glue the 1st layer
② Assembly of the
2nd layer module
③ Install the 2nd layer module
④ Firmly fix the center shaft
pump
⑤ degassing
⑥ release vacuum

## Slide 19: Installation 2
⑦ Assembly of the
3rd layer half modules
⑨ Install the 3rd layer
pump
⑩ degassing
⑪ release vacuum and
add ribs/MBs
⑧ the shaft is released
⑫ Sealing

## Slide 20: Internal cut view